# Intended Audience

This manuscript is intended for biologists and ecologists interested in monitoring abiotic variables at fine spatial scales. It is not written as a specialist electronics or embedded systems engineering paper. The technical architecture is described in sufficient detail to ensure transparency, reproducibility and practical implementation, but the narrative prioritises clarity over technical depth. Engineering concepts are explained in plain language, with acronyms and hardware terminology introduced only where necessary and defined clearly. The manuscript aims to occupy a middle ground between ecology and instrumentation: rigorous enough to be reproducible, but accessible to researchers without formal training in electronics or firmware development. The emphasis is on ecological applicability, system design logic and deployment considerations rather than low-level circuit analysis or firmware optimisation.

# Title

Design and initial testing of a low-cost, open microclimate monitoring system for fine-scale ecological applications

A modular ESP32-based logger with an ultrasonic anemometer and integrated PAR sensing

1. Introduction

Microclimate, the climatic conditions experienced by organisms at spatial scales of centimetres to metres, plays a central role in shaping physiological performance, behaviour and survival. Because organisms respond to the thermal, radiative and moisture conditions of their immediate surroundings rather than to regional atmospheric averages, fine-scale climatic variation often exerts stronger ecological influence than coarse-scale weather data (Kemppinen et al., 2024). Standard meteorological stations are intentionally designed to minimise local variability, measuring temperature, humidity and wind under well-ventilated, standardised conditions at fixed heights above ground. While essential for regional climate monitoring, such measurements frequently fail to capture the heterogeneity experienced within soil layers, vegetation canopies, cavities or near-surface boundary layers (De Frenne et al., 2025; Maclean et al., 2021).

Microclimatic conditions can vary dramatically over short horizontal and vertical distances. Vegetation structure, surface albedo, soil moisture, topography and airflow interact to generate mosaics of temperature, humidity and radiation that change over metres or even tens of centimetres. Near the ground, conductive and convective heat exchange create steep gradients that diverge substantially from free-air meteorological measurements (Maclean et al., 2021). Temporal variability is equally important. Biologically meaningful extremes such as rapid heating events, short-term thermal fluctuations or transient wind shifts can unfold over minutes to hours (Chen et al., 2024; Maclean et al., 2021). Capturing this spatiotemporal heterogeneity is increasingly recognised as essential for predicting ecological responses to climate change, identifying microrefugia and informing habitat management and restoration (De Frenne et al., 2025)

Solar photovoltaic (PV) farms represent a rapidly expanding land use that generates complex microclimatic mosaics (Liu et al., 2026). The arrangement of elevated panels, inter-row spacing and supporting infrastructure creates distinct microhabitat zones, including under-panel environments, open inter-row corridors and transitional edge areas. These zones differ in incident radiation, wind exposure, surface temperature and moisture dynamics as panels modify incoming solar radiation and alter near-surface energy and water fluxes (Li et al., 2025; Liu et al., 2026; Wu et al., 2022). Under-panel areas typically experience substantial reductions in shortwave radiation and moderated soil temperatures (Li et al., 2025; Wu et al., 2022), while inter-row zones remain more strongly coupled to atmospheric forcing. Edge zones can combine shading gradients with lateral redistribution of water and airflow, generating additional spatial heterogeneity (Liu et al., 2026; Wu et al., 2022). Such spatial structure occurs at meter and sub-meter scales and is largely invisible to conventional meteorological monitoring. As solar energy infrastructure expands globally, the spatial complexity of PV-induced microclimates underscores the need for scalable, modular monitoring systems designed specifically to resolve meter-scale environmental gradients within renewable energy landscapes.

Despite growing recognition of the importance of microclimate, a monitoring gap persists. High-end atmospheric systems, including eddy covariance towers and World Meteorological Organization (WMO) style stations, are designed to quantify ecosystem-scale fluxes and spatially integrated atmospheric conditions (Baldocchi, 2003; World Meteorological Organization, 2018). Global networks have demonstrated the power of these approaches for understanding carbon, water and energy exchange across diverse ecosystems, including FLUXNET and long-running regional programs such as OzFlux (Baldocchi et al., 2001; Beringer et al., 2022). Large observatory networks such as NEON further extend this approach through highly standardised tower designs, integrated sensor suites and open data products (Metzger et al., 2019). However, by design, eddy covariance systems integrate processes over footprint scales of tens to hundreds of meters, effectively aggregating fine-scale microclimatic variability into ecosystem-level flux estimates. While recent advances, including camera-based approaches to estimate flux components, have expanded observational capability, these methods still operate at spatial scales that exceed the centimetre- to metre-scale heterogeneity experienced by organisms. Consequently, instrumentation optimised for ecosystem-scale flux integration is not inherently configured to resolve fine-scale microclimatic gradients within structurally heterogeneous environments such as solar farms.

In addition to scale considerations, system architecture introduces further constraints. Sensor combinations are often predefined within specific hardware configurations, limiting flexibility in adapting instrumentation to evolving research questions (Hayibo & Pearce, 2025). In remote deployments, data collection systems frequently require periodic physical access for configuration, data retrieval or maintenance, which becomes increasingly complex as network size grows (Ďud’ák et al., 2023). Long-term autonomous environmental monitoring also imposes constraints on power consumption, storage efficiency and firmware optimisation, necessitating trade-offs in hardware and file system design (Ďud’ák et al., 2023). Importantly, where fine-scale heterogeneity requires dense spatial replication, per-unit hardware cost becomes directly linked to achievable sampling density. Even where individual instruments perform reliably, scaling deployment to tens or hundreds of nodes can constrain study design and limit spatial resolution.

Although open-source platforms have demonstrated that modular systems can reduce per-unit cost while improving flexibility (Ali et al., 2016; Hayibo & Pearce, 2025) ,comparatively fewer studies provide detailed, deployable descriptions that integrate enclosure geometry, long-term field durability, firmware architecture and ecological deployment strategy into cohesive monitoring frameworks. In microclimate research, enclosure design and sensor placement directly influence measurement bias and response time, as airflow, radiation exposure and shielding affect both accuracy and temporal sensitivity. System performance is therefore determined not only by sensor specification, but by the interaction between hardware architecture, mechanical design and deployment context (Maclean et al., 2021). Monitoring platforms intended for structurally heterogeneous solar photovoltaic landscapes must consequently be compact, mechanically stable and capable of resolving fine-scale gradients in radiation, airflow and surface conditions across under-panel, inter-row and edge environments.

In this study, we present the design and implementation of an open, reproducible microclimate monitoring system tailored for fine-scale deployment in heterogeneous landscapes such as photovoltaic (PV) installations. The system integrates temperature and relative humidity sensing, photosynthetically active radiation (PAR) measurement, and ultrasonic wind speed and direction sensing within a compact architecture engineered for field robustness. A central design principle is modularity: sensing modules are decoupled from core data handling and power management, and standard digital communication interfaces (e.g. I²C and 1-Wire) are exposed to enable additional sensors to be incorporated without hardware redesign and with only minor firmware modification. This architecture allows expansion to include variables such as soil moisture, soil temperature, or additional radiation and atmospheric sensors as site-specific research questions evolve, while maintaining a consistent and reproducible platform.

Rather than pursuing comprehensive metrological validation against reference-grade atmospheric instrumentation, this work emphasises transparent documentation of system architecture, electronics design, firmware structure, mechanical housing considerations and build workflow, alongside initial functional testing. The objective is to provide a replicable and adaptable framework that lowers barriers to adoption and supports dense, site-specific microclimate monitoring. By integrating electronics, firmware and enclosure design into a unified and expandable system description, we seek to bridge the gap between proof-of-concept sensor demonstrations and deployable ecological instrumentation.

As renewable energy landscapes continue to expand, scalable tools capable of resolving metre-scale environmental heterogeneity will become increasingly important. Open, modular microclimate systems provide a pathway toward flexible, spatially explicit monitoring of the environments that organisms actually experience.

# 2. System Design Philosophy

The monitoring platform was developed as a fully custom hardware and firmware system guided by five core principles: reproducibility and cost efficiency, distributed node networking with centralised aggregation, modular sensor integration, aerodynamic measurement clarity, and environmental robustness with field configurability.

### 2.1 Reproducible and Cost-Efficient Architecture

The system was implemented using custom-designed printed circuit boards (PCBs) developed in EasyEDA and fabricated via JLCPCB. Component selection prioritised widely available parts to ensure reproducibility and supply stability. Integration of sensing interfaces, power regulation and communication subsystems within compact layouts reduced per-unit cost while maintaining control over signal integrity and environmental performance. This approach supported scalable replication without reliance on proprietary hardware ecosystems.

### 2.2 Distributed Node Network with Centralised Aggregation

The architecture comprised multiple sensor nodes communicating with a central aggregation unit (“mothership”). Nodes function as lightweight sensing units, transmitting measurements wirelessly using the ESP-NOW (WIFI) protocol. Centralised logging at the mothership provides unified data storage and simplified retrieval, reducing direct interaction with individual nodes. This distributed structure supports dense spatial deployment while maintaining manageable data handling and network stability.

### 2.3 Modular Sensor Integration

Modularity was central to the design to accommodate evolving ecological research needs. Standard digital interfaces, including I²C and 1-Wire, were exposed at the hardware level to allow integration of commonly available environmental sensors without redesign of the core architecture. Firmware supports sensor abstraction, enabling new devices to be incorporated through software configuration rather than hardware modification.

The PCB also provides breakout access to multiple analog input channels with selectable excitation voltage levels, enabling compatibility with sensors requiring differing supply and signal conditions. This supports integration of voltage-output and resistive sensors without additional signal-conditioning hardware.

### 2.4 Aerodynamic Measurement Clarity

To preserve measurement accuracy, particularly for ultrasonic wind sensing, the sensing region was designed to minimise geometric obstruction and airflow disturbance. The upper assembly was kept structurally simple to reduce radiative interference and flow distortion that could bias anemometer readings. Electronics, battery systems and wiring were isolated within a lower enclosure to prevent thermal contamination and aerodynamic interference.

### 2.5 Environmental Robustness and Field Configurability

The system was engineered for outdoor deployment, with electronics housed in sealed enclosures to ensure water resistance and mechanical protection. Mechanical components were designed for fused deposition modelling (FDM) 3D printing, allowing rapid fabrication and iterative refinement. Nodes are configured via a local Wi-Fi interface hosted by the mothership, enabling adjustment of measurement intervals, wake schedules and sensor states without physical access or firmware recompilation. This reduces field maintenance burden and supports adaptive deployment strategies.

# 3. System Components (lead with sensors)

## Instrumentation overview

(Frame outputs, not electronics)

- Air temperature & humidity

- Soil Moisture

- Radiation (PAR)

- Wind speed and direction (ultrasonic)

- Pressure - Others

## Logger network (electronics + firmware together)

The logger network is organized around a centralized mothership coordinating multiple distributed sensor nodes for autonomous environmental monitoring under field conditions. The architecture combines low-power duty-cycled operation, synchronized timing, wireless aggregation, and centralized archival in an analysis-ready format.

The mothership serves as network coordinator, time authority, and data logger. A DS3231 real-time clock (RTC) maintains the system reference time used for deployment and timestamping. Wireless communication is implemented using ESP-NOW, a connectionless protocol supported by the ESP32 platform, enabling node discovery, pairing, deployment control, schedule updates, and telemetry reception. Measurements are written to SD-card storage as continuously appended CSV records. A browser-based interface, hosted over a local Wi-Fi access point, provides field configuration of node identity, deployment state, sampling behaviour, and data export.

Each node is an ESP32-based sensing endpoint with a fixed core payload and configurable expansion. The core payload comprises a PAR sensor, an ultrasonic wind anemometer, and an SHTC3 temperature and relative-humidity sensor. In addition, each node provides four auxiliary input channels intended to remain flexible at deployment time. These channels may be configured as analog inputs, I²C interfaces, or mixed-mode combinations within a single node. This approach allows a single hardware platform to support heterogeneous probe sets without redesign.

Node firmware follows a lifecycle model (unpaired, paired, deployed, and stopped/unpaired). During deployment, time and schedule parameters are provisioned by the mothership, after which the node enters periodic acquisition and transmission. Persistent runtime state—including mothership association, deployment status, sampling configuration, and synchronization metadata—is stored in non-volatile memory to support deterministic restart behaviour after power interruption.

At each scheduled acquisition event, the node samples the core sensors and all enabled auxiliary channels, then transmits structured packets to the mothership over ESP-NOW. The mothership timestamps and stores these records together with node and sensor metadata in a unified CSV stream. The resulting system provides a modular, field-deployable framework that integrates synchronized embedded sensing, mixed-interface expansion, wireless telemetry, and centralized data management.

# 3.3 Ultrasonic Anemometer (Electrical Design)

## Time-of-Flight Principle

Wind velocity was measured using a differential time-of-flight (ToF) method. Short ultrasonic bursts (~40 kHz) were transmitted between opposing transducers arranged along orthogonal axes. The propagation time of each pulse was measured in both forward and reverse directions. Wind speed was derived from the difference in travel time (Δt) between opposing paths, while the mean travel time provided an estimate of the effective speed of sound along the acoustic path. This differential approach reduces systematic timing offsets introduced by the electronics and improves robustness to environmental variability such as temperature fluctuations.

## Transmit Path

The transmit (TX) chain generates high-amplitude ultrasonic bursts to ensure adequate acoustic energy across the measurement path. A ~40 kHz carrier signal is generated by the microcontroller using a hardware PWM peripheral. This signal is buffered by a high-speed gate driver, which switches a MOSFET stage powered from a dedicated 22 V boost converter. The resulting voltage swing drives the ultrasonic transducers with sufficient amplitude to overcome acoustic attenuation along the reflector-based measurement path.

Directional excitation of the transducer array is achieved using discrete MOSFET switching stages that connect the selected transducer pair to the transmit driver. These switching elements allow the firmware to sequentially excite orthogonal transducer pairs and to reverse the electrical polarity of the drive signal, enabling bidirectional measurements along each axis.

To reduce electrical noise during the receive phase, the high-voltage transmit supply is enabled only during burst emission. The transmit envelope signal generated by the microcontroller is rectified and filtered using a diode–RC network, producing a smoothed control signal that enables the 22 V boost converter only during transmit bursts. This strategy minimizes switching noise and prevents the high-voltage circuitry from interfering with the sensitive receive electronics.

## Receive Path

The receive (RX) chain converts weak ultrasonic returns into a digital timing signal suitable for precise time measurement. The selected transducer pair is routed through an analog multiplexer that isolates the active receive channel while disconnecting inactive transducers.

The received signal is first passed through an input protection network that limits voltage excursions and provides basic filtering. The signal is then AC-coupled and referenced to a mid-rail virtual ground generated from the system supply. This virtual reference establishes a stable operating point for the subsequent amplification stages.

Two cascaded operational-amplifier stages provide band-limited gain centered near the transmit frequency (~40 kHz). The first stage acts as a low-noise pre-amplifier, while the second stage implements additional active filtering and gain. This band-pass amplification improves signal-to-noise ratio by attenuating low-frequency environmental noise and high-frequency electrical interference.

The conditioned waveform is then fed into a high-speed comparator, which converts the analog ultrasonic waveform into a digital timing signal. The comparator threshold is referenced to the mid-rail voltage, allowing the first valid threshold crossing of the received waveform to be detected reliably. The comparator output is routed directly to a hardware timer input on the microcontroller, enabling precise capture of the arrival time.

## Timing and Processing

Each measurement cycle begins with the generation of a short ultrasonic burst followed by a defined receive window. A brief blanking interval immediately after transmission prevents false triggering caused by transducer ringing or electrical switching artifacts. During the receive window, the first valid comparator edge is captured by a hardware timer and recorded as the time-of-flight measurement.

Measurements are performed sequentially in both directions along each axis. Wind velocity is derived from the difference between the forward and reverse travel times, which removes fixed electronic delays and reduces sensitivity to environmental drift. The mean propagation time is used to estimate the effective speed of sound along the path, providing an additional diagnostic parameter for temperature and measurement health.

Detailed schematics of the transmit and receive circuitry are provided in the Supplementary Materials.

# 3.4 Mechanical design

### The enclosure was developed as a compact, field-ready housing for a custom environmental sensor node designed to measure air temperature and relative humidity, photosynthetically active radiation (PAR), and wind speed and direction using an ultrasonic time-of-flight (ToF) method. In addition to these core measurements, the system supports modular expansion through external sensor ports, allowing additional environmental sensors to be integrated as required for specific monitoring applications. The mechanical design sought to balance two practical constraints: (i) maintaining unobstructed airflow through the wind sensing region to support accurate measurements, and (ii) providing sufficient protected internal volume for the electronics, power supply, and robust cable interfaces suitable for long-term outdoor deployment while supporting flexible sensor integration.

## Head–roof–belly architecture

The housing is organised into three functional components: a head, roof, and belly. The head contains the ultrasonic wind sensing geometry, the roof provides environmental shielding and supports the PAR sensing assembly, and the belly forms a sealed electronics enclosure. This separation keeps cable penetrations and connector hardware away from the wind measurement volume, reducing the likelihood that external protrusions influence local airflow around the acoustic paths. The arrangement also allows the electronics compartment to be opened for servicing without disturbing the wind sensing geometry.

## Overall aerodynamic considerations

The external geometry was designed to minimise turbulence near the ultrasonic measurement region. Smooth external surfaces and gradual transitions were used to reduce sharp flow separation points and to maintain relatively stable airflow conditions through the sensing region. Structural elements (e.g., fasteners and mounting bosses) were positioned outside the primary acoustic pathways where possible. These considerations are important for ultrasonic anemometry because the method relies on precise timing differences between opposing propagation directions; disturbed local flow near the acoustic path can increase variability in time-of-flight estimates.

## Ultrasonic sensing geometry and reflector design

Wind measurements are obtained using ultrasonic time-of-flight between transducers mounted in four angled pods arranged symmetrically around the head. The design uses a compact single-bounce geometry, where an emitted pulse travels upward to the underside of the roof (acting as a reflector) and then returns to the opposing transducer. This approach allows the sensing region to remain relatively open (no struts spanning the measurement volume), while still providing a well-defined and repeatable acoustic path in a compact form factor.

The pod tilt and roof spacing were selected to align the reflected acoustic path so that the nominal reflection occurs near the roof centre and returns to the opposing pod with minimal geometric miss. This symmetry is important because the wind solution is derived from the difference in reciprocal travel times between opposite directions along each axis; maintaining near-identical geometry in both directions reduces systematic offsets and supports stable calibration. The pod arrangement also elevates the acoustic paths above nearby structural features and reduces the likelihood of obstruction by roof mounts or the enclosure rim.

Each pod incorporates a stepped bore that sets transducer seating depth and helps maintain consistent alignment between builds. While sealing features are included for weather resistance, the primary measurement-driven design choice is the repeatability of transducer positioning and the resulting stability of the acoustic path length and reflection geometry.

## Roof and PAR assembly

The roof provides rain protection and houses the PAR sensor and diffuser assembly. A circular diffuser seat and PCB mounting bosses are integrated into the roof to position the PAR optics and sensor board consistently. To avoid introducing additional openings or external cable runs near the wind sensing region, the PAR I²C cable is routed internally through a dedicated channel integrated into one of the roof standoffs, passing from the roof cavity down into the belly where the main electronics are housed.

In addition to its protective role, the underside of the roof acts as the reflective surface required for the single-bounce ultrasonic ToF path, linking the mechanical roof geometry directly to the wind sensing method.

## Belly (electronics enclosure)

The belly encloses the custom node PCB, battery, and internal wiring while maintaining serviceability and weather resistance. A continuous sealing groove around the rim accommodates a silicone cord gasket, providing a repeatable seal while allowing the enclosure to be reopened for maintenance. Cable penetrations for external sensors are implemented on localised flat mounting pads to support standard glands/connectors and to maintain adequate wall thickness around penetrations for mechanical strength.

# 4. Materials and Methods

## 4.1 Bill of materials

(High-level table in paper)
(Full BOM in supplement)

## 4.2 Build and assembly

- Electronics assembly

- Mechanical assembly

- Print orientations

- Sealing steps

## 4.3 Firmware deployment

- Flashing

- Pairing

- Deployment

- Configuration

# 5. System Performance and Initial Testing

(This replaces a traditional “Results” section)

## 5.1 Mechanical performance

- Assembly success

- Fit tolerances

- Sealing behaviour

- Photos

## 5.2 Network and firmware performance

- Node discovery

- Pairing reliability

- Dashboard operation

- Logging stability

## 5.3 Sensor behaviour (qualitative)

- Wind: plausible vectors and variability

- PAR: light response

- Temperature / humidity: expected diurnal patterns

- Example plots:

  - time series

  - wind rose

## 5.4 Issues and refinements

- what didn’t work

- what you changed

- what remains unresolved

# 6. Discussion

## 6.1 Strengths

- Low cost & accessibility

- Modular, open architecture

- Strong mechanical integration

- Analysis-ready outputs

- Clear pathway to CF-compliant data products

## 6.2 Limitations

- Explicitly acknowledge:

  - no formal calibration yet

  - ultrasonic tuning required

  - PAR optics uncharacterised

  - long-term durability untested

## 6.3 Future directions

- Full validation vs HOBO / research-grade instruments

- Power optimisation

- CFD-informed geometry refinement

- Integration into solar farm ecological studies

# 7. Conclusion

- System designed and built

- Functional performance demonstrated

- Open, reproducible foundation established

- Ready for validation and deployment at scale

Supplementary Materials

Detailed schematics

STL files

Firmware repo

Example datasets

R/Python import scripts

Parts sourcing guide

Ali, A. S., Zanzinger, Z., Debose, D., & Stephens, B. (2016). Open Source Building Science Sensors (OSBSS): A low-cost Arduino-based platform for long-term indoor environmental data collection. Building and Environment, 100, 114-126.

Baldocchi, D., Falge, E., Gu, L., Olson, R., Hollinger, D., Running, S., Anthoni, P., Bernhofer, C., Davis, K., & Evans, R. (2001). FLUXNET: A new tool to study the temporal and spatial variability of ecosystem-scale carbon dioxide, water vapor, and energy flux densities. Bulletin of the American Meteorological society, 82(11), 2415-2434.

Baldocchi, D. D. (2003). Assessing the eddy covariance technique for evaluating carbon dioxide exchange rates of ecosystems: past, present and future. Global change biology, 9(4), 479-492.

Beringer, J., Moore, C. E., Cleverly, J., Campbell, D. I., Cleugh, H., De Kauwe, M. G., Kirschbaum, M. U., Griebel, A., Grover, S., & Huete, A. (2022). Bridge to the future: Important lessons from 20 years of ecosystem observations made by the OzFlux network. Global change biology, 28(11), 3489-3514.

Chen, J., North, M., & Franklin, J. F. (2024). The contributions of microclimatic information in advancing ecosystem science. Agricultural and Forest Meteorology, 355, 110105.

De Frenne, P., Beugnon, R., Klinges, D., Lenoir, J., Niittynen, P., Pincebourde, S., Senior, R. A., Aalto, J., Chytrý, K., & Gillingham, P. K. (2025). Ten practical guidelines for microclimate research in terrestrial ecosystems. Methods in Ecology and Evolution, 16(2), 269-294.

Ďud’ák, J., Gašpar, G., Budjač, R., Sládek, I., & Husár, P. (2023). A low-power data logger with simple file system for long-term environmental monitoring in remote areas. IEEE Sensors Journal, 23(24), 31178-31195.

Hayibo, K. S., & Pearce, J. M. (2025). Jericho open resistive data logger: An open-source modular weather station and monitoring system for long-term solar photovoltaic outdoor experimentation. HardwareX, e00720.

Kemppinen, J., Lembrechts, J. J., Van Meerbeek, K., Carnicer, J., Chardon, N. I., Kardol, P., Lenoir, J., Liu, D., Maclean, I., & Pergl, J. (2024). Microclimate, an important part of ecology and biogeography. Global Ecology and Biogeography, 33(6), e13834.

Li, Z., Tao, J., Zhang, J., Wang, Y., Zhong, Y., Ge, Y., Cheng, Z., Yan, B., & Chen, G. (2025). Effect of photovoltaic power plant on the microclimate and soil: A comprehensive analysis based on time series. Energy, 137551.

Liu, W., Lu, M., Pinos, J., Li, Z., Zhao, J., Ma, Y., & Wu, G.-L. (2026). Solar-tracking photovoltaic systems create divergent microhabitats that enhance vegetation recovery in arid sandy lands. Journal of Arid Environments, 232, 105514.

Maclean, I. M., Duffy, J. P., Haesen, S., Govaert, S., De Frenne, P., Vanneste, T., Lenoir, J., Lembrechts, J. J., Rhodes, M. W., & Van Meerbeek, K. (2021). On the measurement of microclimate. Methods in Ecology and Evolution, 12(8), 1397-1410.

Metzger, S., Ayres, E., Durden, D., Florian, C., Lee, R., Lunch, C., Luo, H., Pingintha-Durden, N., Roberti, J. A., & SanClements, M. (2019). From NEON field sites to data portal: A community resource for surface–atmosphere research comes online. Bulletin of the American Meteorological society, 100(11), 2305-2325.

World Meteorological Organization. (2018). Guide to Instruments and Methods of Observation (WMO-No. 8, Issue. W. M. Organization.

Wu, C., Liu, H., Yu, Y., Zhao, W., Liu, J., Yu, H., & Yetemen, O. (2022). Ecohydrological effects of photovoltaic solar farms on soil microclimates and moisture regimes in arid Northwest China: A modeling study. Science of the Total Environment, 802, 149946.